[In the right-hand bottom corner of each page there is a handwritten signature]

[In the header of each page:]

| [logo:] European Funds Smart Growth | [Polish flag] Republic of Poland | [logo:] PARP [Polish Agency for Enterprise Development] | [logo:] European Union European Regional Development Fund |
| --- | --- | --- | --- |

|  | Consolidated text as of 05.02.2023 ……………….………………… (place, date) Appendix No. 1 to Request for Proposal No. 2 of 29/12/2022 PROPOSAL FORM I name and address of the Contracting Party WOŹNIAK MAGDALENA Magdalena Woźniak UL. TARASOWA 4/109 20-819 LUBLIN II Proposal preparation date ………………………………………………………………………………………………… III Name and address of the Contractor Name: ……………………………………………………………………………………… Address: ………………………………………………………...……………………….. KRS (National Court Register No.): …………………………………………… NIP (tax ID No.): ……………………………………………………………….………. REGON (National Official Business Register No.): ………………………………………………………………………………………….……… Tel.: …………………………………………………………………………………….…... E-mail: .………………………………………………………………………….…………. IV Contact details of the person authorised by the Contractor to submit the proposal and contact the Contracting Party: Name: …...……………………………………………………………………..………… Position:…………………………………………………………………………………… Tel.: ……………………………………………………………………………………….… E-mail: ..…………………………………………………………………………………… V Proposal In response to Request for Proposal No. 2 of 29/12/2022 concerning the award of a contract for the delivery of a mobile MRI unit (semi-trailer), in connection with the implementation of the project entitled “Implementation of innovative functional diagnostics of temporomandibular joints supported by numerical simulation of articular disc biomechanics” under sub-measure 3.2.1 Research for the market of the Smart Growth Operational Programme 2014-2020, co-financed by the European Regional Development Fund – I/we hereby submit the following proposal: Price of the contract. Net amount: ………….………... VAT ………………………..………………… Gross amount ………………..………...; Say: ………………………………………… Say: ………………………………………… Say: …………………………………………. Lead time. ……………… days; Say: …………………………… days. COMPLIANCE WITH THE CONTRACTING PARTY’S ASSUMPTIONS I; GENERAL REQUIREMENTS; Criterion met. 1; Mobile MRI unit (semi-trailer), certified by the supplier of the magnetic resonance imaging device GE SIGNATM Voyager 1.5 T- GE Healthcare.; YES NO. 2; Provision of a set of documents to enable confirmation of the certification.; YES NO. 3; Semi-trailer with integrated steel shielding, with Faraday cage installed.; YES NO. 4; The design of the semi-trailer must ensure good quality of MRI, so that the environment does not affect image quality.; YES NO. 5; A mobile unit that meets all road regulations and requirements in all European countries.; YES NO. II; SEMI-TRAILER DESIGN. 1; Semi-trailer made of sandwich panelling with the following features:. 1.1; Closed-cell foam (similar to that used in the aerospace industry).; YES NO. 1.2; Up to 14 layers of materials, including shielding, contained in one continuous panel, seamless side- and top-wall construction that eliminates exposed fasteners.; YES NO. 1.3; Maximum performance in extreme weather conditions, standard from -31°F/-35°C to +113°F/+45°C.; YES NO. 1.4; Insulation value 0.375K.; YES NO. 1.5; Sealing of the mobile unit preventing water entrainment and corrosion.; YES NO. 1.6; Each of the walls is made of a single element.; YES NO. III; INTERIOR. 1; Furniture. 1.1; Furniture made of lightweight multilayer board coated with high-pressure laminate (HPL).; YES NO. 1.2; Cabinet doors made of MDF covered with HPL on both sides.; YES NO. 1.3; Doors executed with push-button locks to guarantee the doors stay closed during transit.; YES NO. 1.4; Doors mounted with drilled-in hinges.; YES NO. 2; Technical compartment. 2.1; The technical area is located at the front side of the Unit, the accessibility to the area is at the front side of the Unit.; YES NO. 2.2; This area consists of two levels and accommodates the power generating unit and the chiller system.; YES NO. 2.3; The area is protected from potential damage by two lateral sides with air supply and exhaust grids.; YES NO. 3; Equipment room. 3.1; The equipment room is located next to the technical department that contains the hard/software for the scanner.; YES NO. 4; Changing room. 4.1; One changing room for patients.; YES NO. 4.2; Equipment: seat, hooks for clothes, mirror, small shelf, staff call system.; YES NO. 5; Control room. 5.1; Control room equipped with lockable upper cabinets with push-locks.; YES NO. 5.2; Ergonomic desk for 2 operators.; YES NO. 5.3; 2 chairs.; YES NO. 6; Entrances. 6.1; Door made of aluminium profiles with an isolated sandwich panel – in the panel there is a window with dimensions of 600 x 400mm, with a privacy glass.; YES NO. 6.2; Entrance 1:; YES NO. 6.2.1; Entrance reachable with hook-on steps with non-slip finish and with an aluminium handrail.; YES NO. 6.2.2.; The door of this entrance can be locked at 90 degrees and 180 degrees.; YES NO. 6.2.3.; There is an outside light over the entrance.; YES NO. 6.2.4.; A foldable platform constructed in a recess porch can optionally be used in combination with a canopy to protect the patients from the weather conditions.; YES NO. 6.3; Entrance 2:; YES NO. 6.3.1; Entrance door with a canopy.; YES NO. 6.3.2; An electrically operated roller shutter is installed behind the canopy.; YES NO. 6.3.3.; Hydraulic lift providing access to the floor level of the body, suitable for stretchers and wheelchairs, equipped with a railing for safety considerations.; YES NO. 7; Patient lift. 7.1; Located directly at the right side of the body, with a canopy, Dhollandia brand.; YES NO. 7.2; Lift power supply: 400 V.; YES NO. 7.3; Lift can be operated with remote control. In case the remote control breaks down, there is a fixed emergency operating button.; YES NO. 7.4; Maximum load 1,500 kg.; YES NO. 7.5; Platform equipped with handrails and a secure gate.; YES. 8; Examination room. 8.1; Fixed separation wall with soundproofing made of protective glass together with a Faraday cage element.; YES NO. 8.2; Glass panels with a roll curtain and integrated blinds mounted to protect patient privacy.; YES NO. 8.3; Ceiling-mounted main panel of the Faraday cage.; YES NO. 8.4; 2 upper cabinets.; YES NO. 8.5; Coil cabinet with roller shutters.; YES NO. 9; Floor. 9.1; Very strong and durable floor cover in all rooms.; YES NO. 9.2; Floor executed with hygienic plinth.; YES NO. 10; Ceiling. 10.1; Aluminium perforated ceiling with ventilation holes.; YES NO. 10.2; White colour .; YES NO. IV; BODYWORK. 1; Dimensions. 1.1; Length: 13,550 mm, external.; YES NO. 1.2; Width: 2,550 mm, external.; YES NO. 1.3; Height: 4,040 mm, from the floor.; YES NO. 2; Walls. 2.1; Insulated medical sandwich panelling with a U value of 0,48 W/m2.K with an integrated steel shielding, according the calculations of the supplier of the MR unit, and with a build-in copper RF shielding; triplex inside and covered with a hard plastic.; YES NO. 2.2; Wall colour: to be chosen by the Contracting Party.; YES NO. 3; Floor. 3.1; Insulated medical sandwich panelling with an integrated steel shielding, according to the calculations of the supplier of the MR unit, and with a build-in copper RF shielding.; YES NO. 3.2; Floor covered with antistatic PVC.; YES NO. 3; Floor colour: to be chosen by the Contracting Party.; YES NO. 4; Roof. 4.1; Insulated medical sandwich panelling with an integrated steel shielding and a copper Faraday cage, triplex inside.; YES NO. 4.2; Non-slip path applied.; YES NO. 4.3; Required safety eyebolt for the roof.; YES NO. 4.4; Roof hatch included.; YES NO. 4.5; Marking of safe-to-walk areas.; YES NO. 5; Outdoor lighting. 5.1; 4 external 230 volt lights.; YES NO. 6; Finishing:. 6.1; The vehicle will be delivered in RAL 9010 white.; YES NO. 7; Technical area (Technical area in the front of the semi-trailer, housing the generator and air-conditioning unit). 7.1; Generator:; YES NO. 7.1.1; 60 KVA soundproofed aggregate installed in the bottom of the technical area or 40 KVA soundproofed aggregate installed in the bottom of the technical area; YES NO. 7.1.2.; Engine – Stage 5 diesel engine, Iveco or similar, 1,500 rpm, 4 cylinders, air-cooled.; YES NO. 7.1.3; Power generating unit: Leroy Somer LSA, 230/400 volt, 50 Hz.; YES NO. 7.2; Accessories:; YES NO. 7.2.1; Generator operation monitoring system.; YES NO. 7.2.2; 3x ampere meter 3x frequency meters and hour counter.; YES NO. 7.2.3; Auto start system/ line power / generator power.; YES NO. 7.3; # The complete power generating set is mounted in an acoustic cabinet so as reduce the noise level to a minimum.; YES # NO. 7.3.1; # The diesel tank is mounted under the generator and has a capacity of 220 litres.; YES # NO. 7.3.2; # An indicator lamp which tells whether the power-generating unit is in operation is mounted on the outside of the bodywork.; YES # NO. 7.4; # A separate alarm light is mounted in the control room to indicate when the generator is running, to avoid scanning when power is supplied only from the generator.; YES # NO. 7.5; Chiller system with fan coil technology, in accordance with the medical guidelines of the medical device manufacturer. Cooling system that meets the requirements in EU countries and in the United Kingdom of Great Britain and Northern Ireland. Refrigerating unit with 4 condenser systems and 4 fans.; YES NO. 7.6; Cabinet in the technical room used for mounting the electrical control panel comprising 3 ampere meters, an Hz meter, 3 volt meters, phase tester (multi meter), lighting controls of entire unit, reversing switch for power generating unit and phase reversing switch including indicator light or Main electrical cabinet mounted in the lower compartment with Socomec device with automatic switching of power from mains to generator power, phase rotation protection and high voltage protection.; YES NO. 7.7; 10 separate 230 volt sockets in various places.; YES NO. 7.8; 230 volt and 400 volt electrical system.; YES NO. 7.9; Unit connected with 250 amp/400 volt Marechal connector.; YES NO. 7.10; A radio to provide music in the MRI room integrated into one of the upper cabinets of the control room.; YES NO. 7.11; NEN1010-certified MRI semi-trailer, prepared also for BS standards.; YES NO. 7.12; Electrical system adapted to both EU countries and the United Kingdom of Great Britain and Northern Ireland.; YES NO. V; TECHNICAL SPECIFICATION OF THE SEMI-TRAILER. 1; Semi-trailer chassis dimensions. 1.1; Dimension; YES NO. 1.1.1; Length: 13,550 mm; YES NO. 1.1.2; Width: 2,550 mm; YES NO. 1.1.3; Neck beam height: 160 mm; YES NO. 1.1.4.; Coupling height: 1,200 mm; YES NO. 1.1.5; Chassis height: 1,360 mm; YES NO. 1.2; Weight; YES NO. 1.2.1.; Coupling pressure: 12,000 kg; YES NO. 1.2.2; Total axle pressure: 20,000 kg; YES NO. 1.2.3; Total admissible load: 32,000 kg; YES NO. 2; Chassis. 2.1; Flat aluminium frame with 2 main beams connected by cross-beams and edge profile.; YES NO. 2.2; Chassis layout based on a two-axle truck.; YES NO. 2.3; Black colour.; YES NO. 3; Pneumatic/Electrical connections and key locker. 3.1; At the front, a cross-beam constructed with a bench for electrical and pneumatic connections (EU standard).; YES NO. 3.2; EBS/ABS connection.; YES NO. 3.3; 24 V connection.; YES NO. 3.4; Electrical connection 1x7 and 1x15 pins.; YES NO. 3.5; Pneumatic connections.; YES NO. 3.6; Key locker for spare keys on front side cross-beam.; YES NO. 4; LED lighting and marking. 4.1; 24 V 7-pole electrical system.; YES NO. 4.2; Rear lights with integrated brake lights and indicator lights.; YES NO. 4.3; 1 fog light.; YES NO. 4.4; 2 reversing lights and license plate lights. Orange side lights.; YES NO. 4.5; 2 white wide-beam lights at the front.; YES NO. 4.6; 2 red wide-beam lights at the rear.; YES NO. 4.7; 2 side indicator lights.; YES NO. 4.8; 2 license plates.; YES NO. 5; # Vehicle preservation. 5.1; # The entire chassis sprayed with a mono-component filling primer, coat thickness of at least 30 microns, and a highly filling synthetic coating, total dry layer thickness of at least 120 microns.; YES # NO. 6; Axles and suspension. 6.1; Axles with air suspension with 2 low-maintenance ECO axles.; YES NO. 6.2; Technical load capacity per axle 10,000 kg with brake drums 420 x 200 mm and asbestos-free brake lining.; YES NO. 6.3; The system is equipped with an ascending and descending valve and automatic brake adjusters.; YES NO. 6.4; Both axles equipped with spring brake cylinders for parking.; YES NO. 6.5; The last axle delivered as a mechanical steered axle.; YES NO. 7; Steering system. 7.1; Mechanical steering on the last axle.; YES NO. 8; Braking system. 8.1; Wabco 2-line braking system connecting in a load-dependent manner to the air system and spring brake system.; YES NO. 8.2; On both axles, Wabco ABS EBS 4S/3M system. D-version.; YES NO. 9; Wheels. 9.1; 8 steel (optionally aluminium) rims, hub centring, with tyres 275/70 R 22.5 XZA.; YES NO. 10; Levelling. 10.1; Mobile MRI trailer equipped with 2 individual hydraulic legs (prepared for temperatures to -25°C) combined with 2 mechanical levelling legs.; YES NO. 10.2; At the rear of the semi-trailer, 2 electrical stands which can be shimmed with composite levelling blocks.; YES NO. 10.3; The supplier will deliver 2 sets of levelling blocks: - 4x (400 x 400 x 40), - 4x (400 x 400 x 80).; YES NO. 11; Storage area in the chassis. 11.1; At the bottom of the vehicle, stainless steel belly lockers with door design according to health and safety regulations.; YES NO. VI; ADDITIONAL EQUIPMENT. 1; Monitoring system. 1.1; Monitoring system including the following sensors: - cooling system temperature, - power supply line connection, - humidity, - position of the trailer, - temperature.; YES NO. 1.2; Fire alarm system that meets the requirements in EU countries and in the United Kingdom of Great Britain and Northern Ireland, prepared to meet the BS standard.; YES NO. 2; Additional accessories inside the unit. 2.1; A wall picture inside the MRI room.; YES NO. 2.2; Coloured LED lighting inside the MRI room.; YES NO. 2.3; Virtual skylight: The ceiling above the MRI bed will be executed with a virtual skylight. Design as arranged with the Contracting Party. Skylight dimensions 1,200 mm x 1,800 mm.; YES NO. 3; Outside livery of the semi-trailer. LIVERY as per the Contracting Party’s design.; YES NO. VII; DELIVERY DATE. 1; Delivery: 30/06/2023; YES NO. VIII; WARRANTY, MAINTENANCE AND REPAIR SERVICE AND QUALITY. 1; Warranty:. 1.1; 60 months for the sandwich panel construction.; YES NO. 1.2; 24 months for materials and parts, excluding wear and tear products.; YES NO. 1.3; The warranty is valid in the European Union, Norway, Switzerland and the United Kingdom of Great Britain and Northern Ireland.; YES NO. 2; Maintenance service. 2.1; Service and maintenance contract for 24 months covering the European Union, Norway, Switzerland and the United Kingdom of Great Britain and Northern Ireland.; YES NO. 2.2; The preventive service and maintenance contract comes into effect on the date of delivery of the mobile MRI unit.; YES NO. 2.3; Preventive service and maintenance every 6 months, the first 6 months after the delivery date of the mobile MRI unit.; YES NO. 2.4; Upgrade of roll curtains with integrated blinds: 2 x RF window and 1 x entrance door window.; YES NO. 3; Life cycle. 3.1; At least 15 years.; YES NO. 4; Quality. 4.1; ISO 9001 certificate; YES NO. IX; PAYMENT TERMS. 1; Stage 1 – 40% of the order value, payable 30 days after order placement Stage 2 – 30% of the order value, payable 60 days after order placement Stage 3 – 20% of the order value, payable upon installation of magnetic resonance Stage 4 – 10% of the order value, payable upon signing the acceptance certificate and before the transport; YES NO. VI Deadline for order performance The order will be performed by the deadline of …………………… VII Validity of the proposal ………………………………………………. VIII Contact to the person supervising the order on the Contractor’s part Full name: …...……………………………………………….………………………… Position:…………………………………………………………………..………………. Tel.: …………………………………………………………………………………………. E-mail: ..…………………………………………………………………………………… IX Statements of the Contractor I/We hereby state that I/we have read the content of the Request for Proposal and its Appendices and I/we do not raise any objections to them. I/We declare that I/we have obtained all the information necessary to prepare a proposal and sign the contract. I/We declare that we consider ourselves bound by the terms of the present proposal for the period specified in the Request for Proposal. I/We have read Appendix 3 to the Request for Proposal concerning Articles 13 and 14 of Regulation (EU) 2016/679 of the European Parliament and of the Council of 27 April 2016 on the protection of natural persons with regard to the processing of personal data and on the free movement of such data, and repealing Directive 95/46/EC. [handwritten signature:] Magdalena Woźniak (stamp and signature of the person authorised to represent the Contractor) |
| --- | --- |

I certify this to be a true and accurate translation of an electronic document drawn up in the Polish language.

Anna Broniarek

Certified translator of English, entered on the list of officially sworn translators and interpreters kept by the Polish Minister of Justice under No. TP/56/17

Translation register entry No. 62/2023

Poznań, 06 February 2023

Delete as appropriate

Delete as appropriate

Delete as appropriate